## Sheet1
Table S4. Comparative summary of clinical-grade AI studies in pathology
Study | AI Task | Cohort size | n centers | External validation | Metrics (AUC, accuracy, calibration) | Failure modes | Regulatory status
Slotman et al., 2023 | AI-assisted histopathological grading and mitotic index | 371 patients | Single center | NO | AUC:Up to 0.94/Balanced Accuracy: Up to 88%/Sensitivity: Up to 87%/Specificity: Up to 90% | Staining/Processing Variability /Single-Institution Bias/Class Imbalance/Magnification | Research only
Challa et al., 2022 | Deep learning model for detecting lymph node metastasis in breast cancer | 168 patients | Single center | NO | Sensitivity: 100%/NPV: 100%/Concordance Rate with Pathologists: ~99% /Time Efficiency: Significantly less time using AI vs. IHC (0.6 vs. 1.0 minutes per slide, P=0.0377) | High False-Positive Rate/Screening-Only Tool/Not Evaluated for Other Features/Preanalytical Variability | Research only
Kindler et al., 2023 | Detect lymph node metastases in CRC specimens as a diagnostic decision support tool. | 758 slides | Single center | NO | AUC: 0.940/Accuracy (on slides with cancer): 0.952/Accuracy (on benign slides): 0.996/Interobserver Agreement:Kappa = 0.94 (excellent) without AI assistance | Single-Institute Setup/Staining and Slide Variation/Limited Pathologist Number/Limited Lesion Recognition | Research only
Zhou et al., 2022 | Automated diagnosis and subtyping of non-melanoma skin cancer | 352 slides | Multi-center | Yes | AUC (one-vs-rest): BCC: 0.92 (MuSCID) vs 0.87 (baseline) SCC-In Situ: 0.87 vs. 0.73 SCC-Invasive: 0.92 vs. 0.82 | Domain shift / stain issues/Template organ limitation/Data leakage risk in SOC | Research only
Cai et al., 2021 | AI for pan-cancer Prognostic Analysis | 492 patients | Multi-center | Yes | C-index for CD3_CT: 0.65 (training), 0.69 (validation)/iAUC for CD3_CT : 0.65 (Training), 0.68 (Validation)/Kaplan-Meier Survival Analysis (Log-rank P-value):Training cohort: P < 0.001;Validation cohort: P = 0.005 | Limited external validation centers/Relatively small validation cohort size /Sample size limitation for subgroup analysis/potential for selection bias | Research only
Sakamoto et al.., 2022 | Collaborative workflow AI/Tumour Cellularity Evaluation | 492 patients | Multi-center | Yes | Sensitivity: 97.1% /Specificity: 87.0%/Accuracy (nuclear recognition): 98.5% | Small training dataset/Scanner variation/Inter-institutional variation /Specimen-type dependency/Algorithm errors | Research only
Gu er al., 2023 | AI for CNS tumor classification (Mitosis Quantification and Detection) | Patients: 12 | Single center | NO | AI Model: Precision 0.920, Sensitivity 0.830/Mitosis Detection:verage Pathologist: Precision 0.750, Sensitivity 0.667 | Domain shift, tumor heterogeneity/ Small cohort size/Limited generalizability of AI/Specimen sampling bias | Research only
BCC – basal cell carcinoma, CRC – colorectal cancer, iAUC – integrated area under the curve, NPV – negative predictive value, SCC – squamous cell carcinoma, SOC – same-organ calibration
This table summarizes major peer-reviewed and regulatory-grade AI studies applied to histopathological diagnosis. Columns include study type, cancer type, cohort size, multicenter participation, external validation status, calibration/robustness metrics, and regulatory notes (FDA SaMD / CE-IVDR). The matrix serves as an evidence synthesis for Section 3.2, enabling readers to evaluate trends in dataset design, validation methodology, and performance benchmarking across clinical AI systems.
Reference
1 | Slotman A, Xu M, Lindale K, Hardy C, Winkowski D, Baird R, et al. Quantitative nuclear grading: an objective, artificial intelligence-facilitated foundation for grading noninvasive papillary urothelial carcinoma. Lab Invest 2023; 103: 100155.
2 | Challa B, Tahir M, Hu Y, Kellough D, Lujan G, Sun S, et al. Artificial intelligence-aided diagnosis of breast cancer lymph node metastasis on histologic slides in a digital workflow. Mod Pathol 2023; 36: 100216.
3 | Kindler C, Elfwing S, Öhrvik J, Nikberg M. A deep neural network-based decision support tool for the detection of lymph node metastases in colorectal cancer specimens. Mod Pathol 2023; 36: 100015.
4 | Zhou Y, Koyuncu C, Lu C, Grobholz R, Katz I, Madabhushi A, et al. Multi-site cross-organ calibrated deep learning (MuSClD): Automated diagnosis of non-melanoma skin cancer. Med Image Anal 2023; 84: 102702.
5 | Cai M, Zhao K, Wu L, Huang Y, Zhao M, Hu Q, et al. Artificial intelligence-based analysis of tumor-infiltrating lymphocyte spatial distribution for colorectal cancer prognosis. Chin Med J (Engl) 2024; 137: 421-430.
6 | Sakamoto T, Furukawa T, Pham HHN, Kuroda K, Tabata K, Kashima Y, et al. A collaborative workflow between pathologists and deep learning for the evaluation of tumour cellularity in lung adenocarcinoma. Histopathology 2022; 81: 758-769.
7 | Gu H, Yang C, Al-Kharouf I, Magaki S, Lakis N, Williams CK, et al. Enhancing mitosis quantification and detection in meningiomas with computational digital pathology. Acta Neuropathol Commun 2024; 12: 7.