## Sheet1
Table S5. Regulatory and compliance matrix – device classes, regulatory pathways, post-market controls, PCCP, cybersecurity, and mapping to data/algorithm/system
Device category | Intended use (examples) | Typical regulatory pathway (US / EU / UK) | Post-market surveillance / PMS expectations | Adaptive AI/change control (PCCP) – required elements | Cybersecurity and networked device expectations | Key requirements mapped to: Data/algorithm/system
Slide scanner (brightfield/fluorescence) | Generate diagnostic-quality WSI for primary diagnosis, remote reporting, image acquisition for downstream AI. | US: FDA 510(k) / De Novo pathway (device classification dependent). EU: IVDR / MDR classification depending on intended use. UK: MHRA conformity assessment. (see refs) | Routine QC logs; calibration records; field safety corrective actions (FSCA); periodic performance checks; reporting of adverse events. (cite device regs) | Firmware changes and autofocus algorithm updates must be pre-specified; non-substantial changes documented; substantial changes require re-validation. PCCP should include: change description, risk assessment, retrospective/ prospective performance plan, rollback criteria | Secure boot, signed firmware, authenticated updates, encrypted network comms (TLS), role-based access, audit logs. Regular vulnerability scanning and patch management | Data: scanner metadata (pixel spacing, objective, firmware, QC logs) must be retained and versioned. Algorithm: autofocus or pre-processing embedded in device treated as software component subject to software lifecycle (IEC 62304). System: device QMS, calibration schedule, change logs. (refs)
Image management system/PACS/IMS (incl. DICOM-WSI) | Store, serve, and provide DICOM-WSI files and metadata to viewers / AI pipelines; archive and retrieval. | US: SaMD considerations; vendor may require 510(k) if intended for diagnosis. EU: IVDR if used in clinical decision making. UK: MHRA guidance | ITU and medical device PMS: audit trails, availability/SLA reporting, data integrity monitoring, backup/restoration tests. | Software updates, indexing or metadata schema changes must have regression test plan. PCCP: dataset re-indexing impact assessment; validation of metadata mapping | Data encryption at rest/in transit, role-based access, logging, secure APIs, identity federation (OAuth2/SAML). Business continuity & disaster recovery | Data: ensure metadata fidelity (DICOM tags, FHIR mappings), checksum and audit trail for WSI files. Algorithm: backend services performing pre-processing must be versioned and validated. System: high availability, access controls, audit logs
AI algorithm – detection/triage (assistive) | Automated prioritization / flagging of slides (non-diagnostic assistive triage). | US: Typically SaMD — could be enforcement discretion or 510(k) depending risk; EU: IVDR classification depends on intended use and risk class. Requirements increase if outputs influence clinical decisions. | Real-world performance monitoring, adverse event reporting, drift detection, post-market clinical follow-up | PCCP should specify model update triggers, retraining frequency, validation datasets, performance acceptance criteria (e.g., no > X% AUC drop), and rollback. Pre-defined monitoring metrics & thresholds | Model artifact integrity checks, authenticated model deployment, telemetry to detect inference anomalies. Data privacy in telemetry (de-identification) | Data: training / validation provenance, metadata linkage (scanner, stain lot) required; data governance for retraining. Algorithm: versioning, explainability, calibration checks, bias monitoring. System: CI/CD for model deployment, monitoring dashboards, logging
AI algorithm – diagnostic/quantitative (making diagnostic claims) | Provide measurements, diagnostic classifications, or companion diagnostic information. | US: SaMD regulatory submission expected (510(k) or De Novo / PMA depending risk). EU: IVDR likely requires conformity assessment (Notified Body) and clinical evidence. UK: MHRA conformity. | Rigorous PMS with planned post-market clinical performance studies, periodic safety update reports (PSUR), change control for adaptive learning, and lifecycle management | PCCP mandatory: pre-specified change types, algorithm performance acceptance margins, prospective validation for deployed changes, human-in-the-loop governance. Extensive documentation and traceability for model evolution. | Highest cybersecurity requirements: model provenance protection, integrity verification, secure telemetry, logging of inference inputs/outputs for audit. | Data: labeled clinical datasets with provenance; audit trail for labels; data versioning and consent records. Algorithm: external validation across scanners/sites; calibration & uncertainty quantification. System: integration with LIS for clinician review; fail-safe default to human review
Integrated DP system (scanner + IMS + AI) | End-to-end diagnostic solution (acquisition → AI → reporting). | US/EU/UK: considered a system medical device; combined regulatory evidence required (device + SaMD + IT) | System level PMS: end-to-end performance metrics, interoperability incidents, network security incidents, combined risk assessment | PCCP must include cross-component change control: hardware firmware, AI model updates, IMS schema changes; re-validation plan for integrated behavior | System security architecture, supply chain risk management, network segmentation, continuous monitoring, incident response | Data: cross-component metadata integrity, synchronized timestamps, audit chain. Algorithm: aggregated performance monitoring, drift detection across sites. System: combined risk management file (RMF), integrated change control
Supportive laboratory equipment (microtome, stainer) | Mechanical/chemical processing of tissue; does not directly output diagnostic image but affects image quality. | Often regulated as medical devices (Class II or equivalent) depending on claims; manufacturer QMS (ISO 13485) | Maintenance logs, calibration records, EQA participation (staining proficiency) | Firmware/control logic updates and protocol changes require validation; PCCP includes validation runs with control tissues | Secure device access control; network isolation for legacy/OT devices; logging of protocol changes | Data: run logs, reagent lot IDs; link to WSI metadata. Algorithm: if device includes embedded algorithm (e.g., staining control), treat as software under IEC 62304. System: maintenance schedule, spare parts traceability